## EDMTemplate
SINGAPORE LAND AUTHORITY
EDM CALIBRATION
ORGANISATION |  |  | Singapore Land Authority
OBSERVER |  |  | Tan Seow Hoon
DATE |  |  | 2013-01-14
[9 empty] | STN | E | Z
TOTAL STATION |  |  |  |  | REFLECTOR |  |  |  | 1 | 0 | 26.531
MODEL |  |  | Topcon GPT 7003 |  | MODEL |  |  | Topcon | 2 | 50.8279 | 20.012
SERIAL NO. |  |  | 390161 |  | NO OF PRISM |  |  | 2 | 3 | 122.7894 | 10.715
CONSTANT (mm) |  |  | 0 |  | CONSTANT (mm) |  |  | 0 | 4 | 215.8487 | 9.183
HT. ABOVE BASEPLATE (m) |  |  | 0.223 |  | HT. ABOVE BASEPLATE (m) |  |  | 0.225 | 5 | 298.3375 | 9.274
[9 empty] | 6 | 359.7298 | 12.123
FROM | TO | OBSERVED SLOPE |  | REDUCED HORZ. |  | KNOWN HORZ. |  | RESIDUAL (m) | 7 | 399.9903 | 17.113
|  | DISTANCE (m) |  | DISTANCE (m) |  | DISTANCE (m)
[9 empty] | dH | Known SD | Y | X | (Y-Ybar) | (X-Xbar) | Yhat | (Y-Yhat)
1 | 2 | 51.243 |  | 50.8269 |  | 50.8279 |  | 0.0008743 | -6.517 | 51.244 | 0.99192 | 0.051243 | 0.83381 | -0.13885 | 0.11762 | 0.0008743
1 | 3 | 123.801 |  | 122.7868 |  | 122.7894 |  | 0.0024127 | -15.814 | 123.8036 | 2.5514 | 0.1238 | 2.3933 | -0.066291 | 0.13878 | 0.0024127
1 | 4 | 216.542 |  | 215.8461 |  | 215.8487 |  | 0.0023908 | -17.346 | 216.5446 | 2.5566 | 0.21654 | 2.3985 | 0.02645 | 0.16582 | 0.0023908
1 | 5 | 298.833 |  | 298.3344 |  | 298.3375 |  | 0.0028839 | -17.255 | 298.8361 | 3.0737 | 0.29883 | 2.9156 | 0.10874 | 0.18982 | 0.0028839
1 | 6 | 360.015 |  | 359.7267 |  | 359.7298 |  | 0.0029333 | -14.406 | 360.0181 | 3.141 | 0.36001 | 2.9829 | 0.16992 | 0.20766 | 0.0029333
1 | 7 | 400.104 |  | 399.9932 |  | 399.9903 |  | -0.0031057 | -9.416 | 400.1011 | -2.8863 | 0.4001 | -3.0444 | 0.21001 | 0.21935 | -0.0031057
2 | 3 | 72.557 |  | 71.9592 |  | 71.9615 |  | 0.0021933 | -9.295 | 72.5593 | 2.3172 | 0.072557 | 2.1591 | -0.11754 | 0.12383 | 0.0021933
2 | 4 | 165.373 |  | 165.0182 |  | 165.0208 |  | 0.0024469 | -10.827 | 165.3756 | 2.5978 | 0.16537 | 2.4397 | -0.024719 | 0.1509 | 0.0024469
2 | 5 | 247.74 |  | 247.5073 |  | 247.5096 |  | 0.0021585 | -10.736 | 247.7423 | 2.3335 | 0.24774 | 2.1754 | 0.057648 | 0.17492 | 0.0021585
2 | 6 | 308.999 |  | 308.8983 |  | 308.9019 |  | 0.0033777 | -7.887 | 309.0026 | 3.5705 | 0.309 | 3.4124 | 0.11891 | 0.19278 | 0.0033777
2 | 7 | 349.178 |  | 349.166 |  | 349.1624 |  | -0.0037865 | -2.897 | 349.1744 | -3.582 | 0.34918 | -3.7401 | 0.15909 | 0.2045 | -0.0037865
3 | 4 | 93.071 |  | 93.0584 |  | 93.0593 |  | 0.0007468 | -1.53 | 93.0719 | 0.87661 | 0.093071 | 0.71851 | -0.097021 | 0.12982 | 0.0007468
3 | 5 | 175.554 |  | 175.5481 |  | 175.5481 |  | -0.00015609 | -1.439 | 175.554 | -0.0022255 | 0.17555 | -0.16033 | -0.014538 | 0.15387 | -0.00015609
3 | 6 | 236.943 |  | 236.9388 |  | 236.9404 |  | 0.0014236 | 1.41 | 236.9446 | 1.5953 | 0.23694 | 1.4372 | 0.046851 | 0.17177 | 0.0014236
3 | 7 | 277.28 |  | 277.2061 |  | 277.2009 |  | -0.0054119 | 6.4 | 277.2748 | -5.2284 | 0.27728 | -5.3865 | 0.087188 | 0.18353 | -0.0054119
4 | 5 | 82.488 |  | 82.4879 |  | 82.4888 |  | 0.0007257 | 0.093 | 82.4889 | 0.85243 | 0.082488 | 0.69432 | -0.1076 | 0.12673 | 0.0007257
4 | 6 | 143.91 |  | 143.8799 |  | 143.8811 |  | 0.0010304 | 2.942 | 143.9112 | 1.175 | 0.14391 | 1.0169 | -0.046182 | 0.14464 | 0.0010304
4 | 7 | 184.317 |  | 184.1462 |  | 184.1416 |  | -0.004798 | 7.932 | 184.3124 | -4.6416 | 0.18432 | -4.7997 | -0.0057753 | 0.15642 | -0.004798
5 | 6 | 61.459 |  | 61.3928 |  | 61.3923 |  | -0.00065739 | 2.851 | 61.4585 | -0.53679 | 0.061459 | -0.6949 | -0.12863 | 0.1206 | -0.00065739
5 | 7 | 101.958 |  | 101.656 |  | 101.6528 |  | -0.0033727 | 7.841 | 101.9548 | -3.2403 | 0.10196 | -3.3984 | -0.088134 | 0.13241 | -0.0033727
6 | 7 | 40.573 |  | 40.2647 |  | 40.2605 |  | -0.0043097 | 4.992 | 40.5688 | -4.1952 | 0.040573 | -4.3533 | -0.14952 | 0.11451 | -0.0043097
[9 empty] | N |  | Ybar | Xbar | SXX | SXY |  | SSE
ADDITIVE CONSTANT (m) | [6 empty] | = | 0.00010268 | 21 | Mean = | 0.15811 | 0.19009 | 0.24621 | 0.071796 |  | 0.00016696
STD DEV OF ADDITIVE CONSTANT (m) | [6 empty] | = | 0.001307 |  |  |  |  | SYY
SCALE (ppm) | [6 empty] | = | 0.29161 |  |  |  |  | 166.9803
STD DEV OF SCALE (ppm) | [6 empty] | = | 5.9742 |  | s = | SXY/SXX | = | 0.29161 | ppm
STD DEV OF SINGLE MEASURED DISTANCE (m) | [6 empty] | = | 0.0029643 |  | a = | Ybar-a*Xbar | = | 0.10268 | mm
[10 empty] | R^2 = | (SYY-SSE*1000000)/SYY |  |  | = | 0.00012538